PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

- It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

- The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

- The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

- PRODUCT {Test} NAME

- INTENDED USE and TEST TYPE (qualitative or qualitative)

- SUMMARY AND EXPLANATION

- PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

- REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

- INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

- SAMPLE COLLECTION/HANDLING

- PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

Free and Total Protein S

ELISA Kit

Cat. No. 5292

INTENDED USE

The Free and Total Protein S Kit is an enzyme-linked immunosorbent assay (ELISA) for the quantitative determination of Free and Total Protein S Antigen in citrated human plasma.

SUMMARY

Protein S is a vitamin K-dependent protein synthesized in the liver, vascular endothelium, megakaryocytes, and with Protein C plays an important physiologic role in the Protein C Anticoagulant System.1, 2 This anticoagulant system is one of the major regulators of hemostasis by inhibiting clot formation and by promoting fibrinolysis. Protein S functions as a cofactor for activated Protein C on the vascular membrane to facilitate the degradation of clotting factors Va and VIIIa. In normal plasma approximately 40% of Protein S circulates as free molecule, while 60% is complexed with C4b, a plasma protein of the classical complement pathway.3 Only Free Protein S is functionally active and able to bind to activated Protein C, while the complexed form of Protein S is not.4

Protein S deficiency, either congenital or acquired, may lead to serious thrombotic events such as thrombophlebitis, deep vein thrombosis, or pulmonary embolism. The prevalence of Protein S deficiency has been estimated to be less than 1 case per 300 in the general population. Two-thirds of patients with a congenital deficiency of Protein S levels (less than 50%) may present with venous thrombosis in young adulthood.5, 6 In young patients (< 35 years) with a history of thrombosis, the prevalence may be as high as 15 to 18%.7 Acquired Protein S deficiency may be seen during pregnancy, oral contraceptives, liver disease, oral anticoagulant therapy, diabetes mellitus, postoperative complications, septicemia and various inflammatory syndromes.8 A decreased Protein S activity in plasma may be the result of low concentrations or abnormal function of the Protein S molecule.

The laboratory diagnosis of Protein S deficiency may require both quantitative and qualitative (functional) determinations. Quantitative determinations of Protein S Antigen are based on immunologic procedures such as radial immunodiffusion in gel, Laurell Rocket immunoelectrophoresis and enzyme-linked immunosorbent assay (ELISA).9, 10 Measurement of plasma level for both Free and Total Protein S are useful to determine the type of defect in patients with Protein S deficiency. ELISA procedures are less labor intensive and offer several advantages including more objective, accurate and reproducible results. In addition, ELISA allows automation with commonly available laboratory instrumentation.

PRINCIPLE

The Protein S Antigen assay is a sandwich ELISA. A capture antibody specific for human Protein S is coated to 96-microwell polystyrene plates. Diluted patient plasma is incubated in the wells allowing any available Protein S to bind to the anti-human Protein S antibody on the microwell surface. The plates are washed to remove any unbound plasma molecules. Bound Protein S is quantitated using a horseradish peroxidase (HRP) conjugated anti-human Protein S detection antibody. Any unbound conjugated anti-human Protein S is washed away after an incubation period. A chromogenic substrate of tetramethylbenzidine (TMB) and hydrogen peroxide (H2O2) is added to develop a colored reaction. The intensity of the color is measured spectrophotometrically at 450 nm in optical density (OD) units. Protein S relative percent concentration in patient plasma is determined against a curve prepared from a reference plasma provided with the kit.

Results obtained from diluted plasma samples not pretreated with polyethylene glycol (PEG) represent the Total Protein S concentration for that sample. To measure Free Protein S, PEG is added to plasma samples prior to beginning the assay to precipitate the Protein S-C4b binding protein complex. The supernatant fraction containing Free Protein S may be tested along with the untreated plasma sample. Both Total (untreated) and Free (PEG-treated) Protein S concentrations are determined following the same procedure as described above using separate reference curves.

REAGENTS

WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY.

1.	Protein S Antigen Microwells		
Ingredients: 96 stabilized antibody coated microwells (12 strips of breakaway wells), with frame holder. Wells are coated with anti-human Protein S antibody.

Preparation for Use: The microwells are ready for use as packaged.

Storage and Stability: Store at 2-8°C. Do not freeze. Microwells are stable until the expiration date indicated on the package.

Signs of Deterioration: Avoid contamination.

2.	Sample Diluent

Ingredients: A blue-green solution containing buffers, salts, and sodium azide as a preservative.

WARNING: DO NOT INGEST. To prevent the formation of toxic vapors, sodium azide should not be mixed with acidic solutions. When discarding reagents containing sodium azide, always flush sink with copious quantities of water. This will prevent the formation of metallic azides which, when highly concentrated in metal plumbing, are potentially explosive. In addition to purging drain pipes with water, plumbing should occasionally be decontaminated with 10% NaOH.

Preparation for Use: The diluent is ready for use as packaged.

Storage and Stability: Store at 2-8°C. The diluent is stable until the expiration date indicated on the package.

Signs of Deterioration: Discard if product shows signs of microbial growth.

3.	ELISA Reference Plasma

Ingredients: Contains human plasma.

WARNING: DO NOT INGEST. Plasma has been tested and shown to be negative for Hepatitis B Antigen (HbsAg) and HIV-1 antibody; however, the plasma should be handled as if capable of transmitting infection.

Preparation for Use: Reconstitute Reference Plasma by adding 0.5 mL deionized water. Swirl gently to mix. Allow to stand for 10 minutes before use for complete dissolution.

Storage and Stability: When stored at 2-8°C, the Reference Plasma is stable until the expiration date indicated on the package. Reconstituted solution is stable for 8 hours when stored at 2-8°C.

Signs of Deterioration: Unreconstituted Reference Plasma should appear as a light yellow, dry plug.

4.	Protein S Conjugate Solution

Ingredients: The red solution contains antibodies, specific for Protein S which have been conjugated with horseradish peroxidase.

WARNING: DO NOT INGEST.

Preparation for Use: The conjugate solution is ready for use as packaged.

Storage and Stability: When stored at 2-8°C, the solution is stable until the expiration date indicated on the package.

Signs of Deterioration: Discard if product shows signs of microbial growth.

5. 	Substrate

Ingredients: Substrate contains 3,3’,5,5’-tetramethyl-benzidine and hydrogen peroxide.

WARNING: IRRITANT, DO NOT PIPETTE BY MOUTH. DO NOT INGEST -  The Substrate can cause irritation to the eyes and skin. Absorption through the skin is possible. 	
Preparation for Use: The Substrate is ready for use as packaged.

Storage and Stability: When stored at 2-8°C, the substrate is stable until the expiration date indicated on the package.

Signs of Deterioration: Substrate should be clear and almost colorless.

6.	Stopping Solution

Ingredients: The solution is 0.36 N Sulfuric Acid.

WARNING: DO NOT INGEST, IRRITANT. DO NOT PIPETTE BY MOUTH. Avoid contact with skin or clothing.

Preparation for Use: Solution is ready for use as packaged.

Storage and Stability: The solution should be stored at 2-8°C and is stable until the expiration date indicated on the package.

7.	Phosphate Buffered Saline Concentrate (PBS)

Ingredients: 33X Phosphate Buffered Saline with 0.01% Tween 20.

WARNING: DO NOT INGEST.

Preparation for Use: Dilute 30 mL PBS Concentrate to 1 liter with deionized water. The pH of the final solution should be 7.4 ± 0.1.

Storage and Stability: When stored at 2-8°C, the PBS is stable until the expiration date indicated on the package. The diluted PBS is stable for 1 year stored at 2-8°C.

Signs of Deterioration: Discard if it shows signs of microbial or cross-contamination.

8.	Free Protein S Reagent

Ingredients: 25% Polyethylene Glycol (PEG).

WARNING: DO NOT INGEST.

Preparation for Use: The reagent is ready for use as packaged.

Storage and Stability: When stored at 2-8°C, the reagent is stable until the expiration date.

INSTRUMENTS

A spectrophotometer capable of reading microwell plates at 450 nm is required.

SPECIMEN COLLECTION AND PREPARATION

Specimen: The plasma collected by venipuncture with either 3.2% or 3.8% sodium citrate as an anticoagulant should be used. Centrifuge sample immediately and remove the plasma.

Storage and Stability: Store at 2-8°C until testing can be performed. If not tested within 1 hour of collection, the sample must be stored at -70°C and tested within 1 month.

PROCEDURE

Materials Provided: The following materials needed for the procedure are contained in the kit.

Protein S:Ag Microwells (96)

ELISA Reference Plasma (3 x 0.5 mL)

Protein S Conjugate Solution (1 x 12 mL)

Sample Diluent (1 x 60 mL)

Substrate Solution (1 x 13 mL)

Stopping Solution (1 x 15 mL)

Phosphate Buffered Saline (1 x 30 mL)

Free Protein S Reagent (1 x 2 mL)

Materials Required but not Supplied:

Specialty Assayed Control 1 (S.A.C. 1) - 5301

Specialty Assayed Control 2 (S.A.C. 2) - 5302

Deionized water

Graduated cylinders

Pipettors (5 and 1000 µL)

Plastic squeeze bottle

Plate reading spectrophotometer capable of reading absorbance at 450 nm

Multichannel pipettors capable of delivering to 8 wells Centrifuge

Procedural Notes

1.	Bring plasma samples and kit reagents to room temperature (15-30°C) and mix well before using, avoid foaming. Return all unused samples and reagents to refrigerated storage as soon as possible.

2.	All dilutions of reference plasma, control, and test plasma must be made just prior to use in the assay.

3.	A single water blank well should be set up on each plate with each run. No sample or kit reagents are to be added to this well. Instead, add 200 µL of deionized water to the well immediately prior to reading the plate in the spectrophotometer. The plate reader should be programmed to “zero” or “blank” against this water well.

4.	Good washing technique is critical for optimal performance of the assay. Adequate washing is best accomplished by directing a forceful stream of wash solution from a plastic squeeze bottle with a wide tip into the bottom of the microwells. An automated microtiter plate washing system can also be used.

5.	Important: Failure to adequately remove residual PBS can cause inconsistent color development of the substrate solution.

6.	Use a multichannel pipettor capable of delivering to 8 wells simultaneously when possible. This speeds the process and allows for more uniform incubation and reaction times for all wells.

7.	Carefully controlled timing of all steps is critical. All dilutions for curve points and samples must be added within a five minute period. Batch size of samples should not be larger than the amount that can be added within this time period.

8.	For all incubations, the start of the incubation period begins with the completion of reagent or sample addition.

9.	Addition of all samples and reagents should be performed at the same rate and in the same sequence.

10.	Incubation temperatures above or below normal room temperature (15 to 30°C) may contribute to inaccurate results.

11.	Avoid microbial and cross-contamination of reagents when opening and removing aliquots from the primary vials.

12.	Do not use kit components beyond expiration date.

13.	Do not use kit components from different kit lot numbers.

STEP-BY-STEP METHOD

1.	Remove any microwell strips that will not be used from the frame holder and store them in the plastic pouch.

2.	Assay each reference plasma dilution in duplicate for Free and Total Protein S. It is advised that duplicate determinations be made for all samples. One well should be run as a reagent blank; sample diluent without plasma is added to the well as explained in step 8 of this section. This well will be treated the same as a control or patient sample in subsequent assay steps. A water blank well should be included with each plate; it is to remain empty until 200 µL of deionized water is added at the completion of the assay, immediately prior to reading the plate. The water blank well is to be used to zero the plate reader.

3.	Pretreatment for Free Protein S determination: All plasma samples to be tested for Free Protein S must be pretreated. Do not dilute plasma samples before PEG pretreatment. Add 15 µL of Free Protein S Reagent (PEG) to 85 µL each control and patient plasma. Add 45 µL PEG to 255 µL Reference Plasma to yield sufficient supernatant for the reference curve. Vortex and place on ice for 30 minutes. Centrifuge for 10 minutes at 3,000 rpm, and use supernatant to prepare curve and sample dilutions as described below.

4.	Predilution of plasmas for Total Protein S deter-mination. For each reference plasma, control plasma, and patient plasma to be tested for Total Protein S a 1:2 predilution with Sample Diluent (100 µL plasma + 100 µL diluent) must be made. These predilutions are then utilized in preparing the working dilutions for Total Protein S in steps 5 and 6.

5.	Using the Reference Plasma provided with the kit, prepare six reference dilutions as described below. (Prepare one set of dilutions with untreated plasma which has been prediluted 1:2 for Total Protein S and a second set with PEG-treated plasma for Free Protein S determination):

Volume Reference	Volume Sample	*Reference

Plasma	Diluent	Level

30 µL	+	500 µL	=	150

20 µL	+	500 µL	=	100

15 µL	+	500 µL	=	75

10 µL	+	500 µL	=	50

10 µL	+	1000 µL	=	25

10 µL	+	2000 µL	=	12.5

* Reference level value to be used for constructing reference curve only.

6.	Prepare a 1:26 dilution of patient sample and control plasma (prediluted 1:2 for Total or PEG treated for Free) in Sample Diluent (blue-green solution); e.g. 20 µL sample added to 500 µL Sample Diluent = 1:26 dilution. Mix thoroughly.

7.	Add 100 µL of the dilutions (reference plasmas x 6, patient samples and controls) to the appropriate microwells for Total and/or Free Protein S determinations.

8.	Add 100 µL of Sample Diluent to the reagent blank well. Place nothing in the well intended for the water blank.

9.	Incubate 40 minutes at room temperature. After the incubation is complete, carefully invert the microwells and decant the sample fluid. Take care to prevent sample from one microwell to flow into another.

10.	Wash 4 times with working PBS solution. Each well should be filled with PBS solution per wash. PBS in the empty water blank well will not interfere with the procedure. Invert microwells between each wash to empty fluid. Use a snapping motion of the wrist to shake the liquid from the wells. The frame must be squeezed at the center on the top and bottom to retain microwell modules during washing. Blot on absorbent paper to remove residual wash fluid. Do not allow wells to dry out between steps.

11.	Add 100 µL Protein S Conjugate Solution (red) to each well (except for the water blank well).

12.	Incubate for 10 minutes at room temperature. After the incubation is complete, carefully invert the microwells and decant the Conjugate Solution.

13.	Wash 4 times with working PBS solution as in step 10. Use a snapping motion to drain the liquid, and blot on absorbent towels after the final wash. Do not allow the wells to dry out.

14.	Add 100 µL Substrate to each well (except for the water blank well) and incubate for 10 minutes at room temperature. Add the substrate to the wells at a steady rate. The substrate will turn blue in wells with positive samples.

15.	Add 100 µL of the Stopping Solution (0.36 N sulfuric acid) to each well (except for the water blank well) to stop the enzyme reaction. Be sure to add the acid to the wells in the same order and at the same rate as the working Substrate Solution was added to the wells. Blue Substrate will turn yellow and colorless substrate will remain colorless. Do not add Stopping Solution to the water blank well. Instead, add 200 µL of deionized water to the water blank well. Blank or zero the plate reader against the water blank well. Read the O.D. of each well at 450 nm. For best results, the O.D. values should be measured immediately after the addition of acid.

Quality Control

1.	The mean O.D. of the reagent blank should be less than 0.1 when the spectrophotometer has been blanked against the water well. Readings greater than 0.1 may indicate possible reagent contamination or inadequate plate washing.

2.	O.D. for the duplicates of the reference plasma dilutions, plasma controls or patient samples should be within 20% of the mean O.D. value for absorbance readings greater than 0.200.

3.	Protein S Antigen values (Free and Total) obtained for the controls should be within manufacturer’s assayed ELISA ranges.

RESULTS

1.	Calculate the mean O.D. for the duplicates of the reference plasma dilutions, controls, and patient samples.

2.	Plot the mean O.D. obtained for each dilution of the reference plasma against the corresponding value of the reference level. The curve may be plotted on a semi-log (if semi-log, plot O.D. on linear axis) or log-log graph.  Draw a line to connect the points. Prepare separate curves for Free and Total Protein S determinations.

3.	Using the mean O.D., determine the control and patient relative values from the graph, or, alternatively, calculate the linear regression for the reference curve. To calculate Free and Total Protein S Antigen levels in % of normal, multiply control and patient relative values obtained from the appropriate reference curve by the corresponding assigned value for the ELISA Reference Plasma.

For example*:	Patient relative value (from the reference curve): 40

Reference plasma assigned value: 105% of normal

Actual patient Protein S Antigen value (as % of normal):  40 x 1.05 = 42%

* Example applies to both Free and Total Protein S calculations.

4.	Ensure that all quality control parameters have been met (see Quality Control) before reporting test results.

REFERENCE RANGES

Free and Total Protein S values are expressed in relative percents as compared to pooled normal plasma.

Range

Total Protein S		60%-150%

Free Protein S		50%-130%

These Protein S ranges are consistent with those published literature and reported by available commercial kits. 6, 10

Each laboratory should periodically determine their own reference range for this assay.

PERFORMANCE CHARACTERISTICS

Detection range:

The detection range for Protein S Antigen assay (Free and Total) is 5-150%. However, the effective range of each run will depend on the assayed value of the reference plasma. For greatest accuracy, samples which generate absorbance readings outside the OD range of the reference curve should be retested at an appropriate dilution.

Precision

Intra-assay:

To determine variability within a plate, three plasma samples with known Protein S levels (one each high, medium, and low) were tested in 16 wells by two operators, on six plates from each of three lots. The data, presented in the following table, shows a mean CV of 10.1% for Total Protein S, and 6.6% for Free Protein S across three lots. In addition, ninety-nine (99) patient samples with Free and Total Protein S levels spanning the entire detection range of the assay were tested in duplicate across 3 lots to demonstrate precision end users may expect when performing the assay according to package insert instructions. As shown in the table, the overall mean CV for duplicates was 7.8% for Total Protein S and 4.7% for Free Protein S.

Inter-assay:

Six (6) commercially prepared, assayed plasma samples with values ranging from 57-159% were tested in duplicate on three lots to determine assay precision between lots. The mean inter-assay CV was 11% for Total Protein S and 10.5% for Free Protein S as seen in the table.

Linearity

Serial two-fold dilutions of Protein S reference plasma samples tested on 3 lots of Helena Protein S Antigen assay demonstrated curves with a mean coefficient of determination (r-squared) of 0.985 for Total Protein S and 0.992 for Free Protein S..

Accuracy

Accuracy was determined by testing mixtures of Protein S reference plasma with predetermined values on Helena Protein S Antigen assay to assess the recovery of their theoretical values. The overall mean percent recovery across 3 lots was for Total Protein S 101% and Free Protein S 98%, with an average variation of 7.1% and 5.0% respectively.

LIMITATIONS OF THE TEST

The Protein S concentration values obtained from this assay are an aid to diagnosis only. Each physician must interpret these results in light of the patient’s history, physical findings, and other diagnostic procedures. Patients with congenital homozygous deficiency of Protein S are rare and may show undetectable levels of Protein S, while those with heterozygous deficiency typically have levels below 50% of normal. Acquired Protein S deficiency may be seen in numerous clinical conditions: neonates (show 20-35% lower levels than adults), liver diseases, diabetes mellitus, pregnancy, oral contraceptive, or anticoagulant therapy and disseminated intravascular coagulation (DIC). Increased levels of Protein S may be seen in patients with  nephrotic syndrome and acute inflammation.5-10 Plasma samples could be inadvertently depleted or degraded of Protein S by improper collection or laboratory processing

REFERENCES

1.	DiScipio, RG, Davie, EW,  Characterization of Protein S; a gamma-carboxyglutamic acid containing protein from bovine and human plasma.  Biochemistry 18:899-904, 1979.

2.	Dahlback, B.  Purification of human vitamin K-dependent Protein 2 and its limited proteolysis by thrombin.  Biochem J 209:837- 846, 1983.

3.	Walker, FJ.  Protein S and the Regulation of Activated Protein C.  Semin Thromb Hemost 10:131-138, 1984.

4.	Dahlback, B.  Protein S and C4b-Binding Protein:  Components involved in the regulation of the Protein C Anticoagulant System.  Thromb Haemost 66:49-61, 1991.

5.	Comp, PC, Nixon, RR, et al.  Familial Protein S deficiency is associated with recurrent thrombosis.  J Clin Invest 74:2082- 2088, 1984.

6.	Walker, FJ.  Protein S and Thrombotic Disease.  PSEBM 200:285- 295, 1992.

7.	Gladson, CL, et al.  The frequency of type I heterozygous Protein S and Protein C deficiency in 141 unrelated young patients with venous thrombosis.  Thromb Haemost 59:18-22, 1988.

8.	Alving, BM, Comp, PC.  Recent advances in understanding clotting and evaluating patients with recurrent thrombosis.  Am J Obstet Gynecol 167:1184-1191, 1992.

9.	Deutz-Terlouw, PP, Ballering, L, et al.  Two ELISA’s for measurement of Protein S, and their use in the laboratory diagnosis of Protein S deficiency.  Clin Chimica Acta 186:321- 334, 1989.

10.	Comp, PC.  Laboratory evaluation of Protein S status.  Semin Thromb Hemost 16:177-181, 1990.

Helena’s ELISA Kits

Cat. No.

von Willebrand Factor ELISA Kit	5290

Protein C ELISA Kit	5291

Free and Total Protein S ELISA Kit	5292

For Sales, Technical and Order Information and Service Assistance, call 800-231-5663 toll free.

Helena Laboratories warrants its products to meet our published specifications and to be free from defects in materials and workmanship. Helena’s liability under this contract or otherwise shall be limited to replacement or refund of any amount not to exceed the purchase price attributable to the goods as to which such claim is made. These alternatives shall be buyer’s exclusive remedies.

In no case will Helena Laboratories be liable for consequential damages even if Helena has been advised as to the possibility of such damages.

The foregoing warranties are in lieu of all warranties expressed or implied including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose.

Pro. 21

9/11(4)