PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

- It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

- The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

- The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

- PRODUCT {Test} NAME

- INTENDED USE and TEST TYPE (qualitative or qualitative)

- SUMMARY AND EXPLANATION

- PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

- REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

- INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

- SAMPLE COLLECTION/HANDLING

- PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

QuickGel® Immuno-Fix Procedure

Cat. No. 3351, 3551, 3551T

INTENDED USE

QuickGel ImmunoFix is intended for the qualitative identification of monoclonal gammopathies in serum, cerebrospinal fluid (CSF) or urine using protein electrophoresis and immunofixation on the SPIFE 3000 or the QuickGel Chamber systems.

Cat. No. 3351 - SPIFE Blade Assembly Application

Cat. No. 3551 - QuickGel Blade Application

Cat. No. 3551T - QuickGel Template Application

For In Vitro Diagnostic use.

Rx Only

SUMMARY

Immunofixation electrophoresis (IFE) is a two stage procedure using agarose gel high resolution electrophoresis in the first stage and immunoprecipitation in the second. There are numerous applications for IFE in research, forensic medicine, genetic studies and clinical laboratory procedures. The greatest demand for IFE is in the clinical laboratory where it is primarily used for the detection of monoclonal gammopathies. A monoclonal gammopathy is a primary disease state in which a single clone of plasma cells produces elevated levels of an immunoglobulin of a single class and type. Such immunoglobulins are referred to as monoclonal proteins, M-proteins or paraproteins. Their presence may be of a benign nature or of uncertain significance. In some cases they are indicative of a malignancy such as multiple myeloma or Waldenstrom’s macroglobulinemia. Differentiation must be made between polyclonal and monoclonal gammopathies because polyclonal gammopathies are only a secondary disease state due to clinical disorders such as chronic liver diseases, collagen disorders, rheumatoid arthritis and chronic infections.

Alfonso first described immunofixation in the literature in 1964.1 Alper and Johnson published a more practical procedure in 1969 as a result of their work devoted to the detection of genetic polymorphisms of ceruloplasmin and Gc-globulin and the conversion of C3 during activation.2 They later extended their studies to genetic polymorphisms of complement components and the identification of alpha1 antitrypsin.3,4 Immunofixation has been used as a procedure for the study of immunoglobulins since 1976.5, 6

The QuickGel IFE methods offer many advantages. These include ease of interpretation, excellent resolution, reagent conservation and rapid turnaround.

PRINCIPLE

Proteins are first resolved by electrophoresis. In the second stage, the soluble antigen and antibody are allowed to react. The resultant antigen-antibody complex(es) may become insoluble (as long as the antibody is in slight excess or near equivalency) and precipitate. The precipitation rate depends on the proportions of the reactants, temperature, salt concentration and the pH of the solution. The unreacted proteins are removed by a washing step and the antigen-antibody complex (which might be visible as a white cloudy band in the unstained gel against a dark background) is visualized by staining. The bands in the protein separation are compared with the precipitin bands obtained with immunofixation.

REAGENTS

1.	QuickGel IFE Gel	
Ingredients: Each gel contains agarose in tris-barbital/MOPS buffer with a stabilizer and a preservative added. 	
WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. CAUTION: DO NOT INGEST. The gel contains barbital which, in sufficient quantity, can be toxic.

Preparation for Use: The gels are ready for use as packaged.	
Storage and Stability: The gels should be stored horizontally at room temperature (15 to 30°C) and are stable until the expiration date indicated on the package. The gels must be stored in the protective packaging in which they are shipped. DO NOT REFRIGERATE OR FREEZE.	
Signs of Deterioration: Any of the following conditions may indicate deterioration of the gel: (1) crystalline appearance indicating the agarose has been frozen, (2) cracking and peeling indicating drying of the agarose, (3) bacterial growth indicating contamination, (4) thinning of gel blocks.

2.	Acid Violet Stain	
Ingredients: The stain is comprised of Acid Violet stain.	
WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. DO NOT INGEST.	
Preparation for Use: Dissolve the dry stain in 1 liter of 10% acetic acid and mix thoroughly. Fill the SPIFE stain vat.	
Storage and Stability: The dry stain should be stored at 15 to 30°C and is stable until the expiration date indicated on the package. The stain solution is stable for six months when stored at 15 to 30°C in a closed container.	
Signs of Deterioration: The diluted stain should be a homogeneous mixture free of precipitate.

3.	Citric Acid Destain	
Ingredients: After dissolution, the destain contains 0.3% (w/v) citric acid.	
WARNING: FOR IN-VITRO DIAGNOSTIC USE. DO NOT INGEST - IRRITANT.	
Preparation for Use: Pour 11 L of deionized water into the Destain vat. Add the entire package of Destain. Mix well until completely dissolved.	
Storage and Stability: Store the Destain at 15 to 30°C. It is stable until the expiration date on the package.

Signs of Deterioration: Discard if solution becomes cloudy.

4.	Tris-Buffered Saline	
Ingredients: The powder contains a Tris base with Tris HCl and sodium chloride.

]WARNING: FOR IN-VITRO DIAGNOSTIC USE

Preparation for Use: Dissolve the powder in 8 L of deionized water and mix thoroughly.	
Storage and Stability: Store the dry powder at 15 to 30°C until the expiration date indicated on the label. The buffer solution should be stored at 15 to 30°C.	
Signs of Deterioration: The buffer solution should be discarded if it shows signs of bacterial contamination.

5.	QuickGel IFE Protein Fixative	
Ingredients: The fixative contains 4.0% sulfosalicylic acid, 6.7%  trichloroacetic acid, 0.2% glutaraldehyde and 1.7% guanidine HCl.	
WARNING: FOR IN-VITRO DIAGNOSTIC USE. CORROSIVE - NEVER PIPETTE BY MOUTH. DO NOT INGEST.

Preparation for Use: The fixative is ready for use as packaged. 
Storage and Stability: The fixative should be stored at 2 to 8°C and is stable until the expiration date indicated on the vial.

Signs of Deterioration: The fixative should be a clear solution.

6. Antisera to Human IgG, IgA, IgM, Kappa Light Chain and Lambda Light Chain	
Ingredients: Antisera vials in the kit contain monospecific antisera to human immunoglobulin heavy chains, IgG, IgM, IgA and to human light chains, Kappa and Lambda. The antisera have been prepared in goat. Each vial of antiserum contains a stabilizer and sodium azide as a preservative.	
WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. To prevent the formation of toxic vapors, do not mix with acidic solutions. When discarding, always flush sink with copious amounts of water. This will prevent the formation of metallic azides which, when highly concentrated in metal plumbing, are potentially explosive. In addition to purging pipes with water, plumbing should occasionally be decontaminated with 10% NaOH.

Preparation for Use: The antisera are ready for use as packaged.	
Storage and Stability: The antisera should be stored at 2 to 8°C and are stable until the expiration date indicated on the vial.	
Signs of Deterioration: Extremely cloudy antisera may be indicative of bacterial contamination.

INSTRUMENT

A SPIFE 3000 Analyzer or the QuickGel Chamber must be used to electrophorese, stain, destain and then dry the gels. Refer to the Operator’s Manual for detailed instructions.

SPECIMEN COLLECTION AND HANDLING

Specimen: Fresh serum, CSF or urine is the specimen of choice.

Interfering Factors:

1. Evaporation of uncovered specimens may cause inaccurate results.

2. Plasma should not be used because the fibrinogen may adhere to the gel matrix resulting in a band in all patterns across the gel.

Storage and Stability: If storage is necessary, samples may be stored covered at 2 to 8°C for up to 72 hours.

PROCEDURE FOR SPIFE 3000

Materials Provided: The following materials needed for the procedure are contained in the QuickGel IFE Kit (Cat. No. 3351). Individual items are not available.

QuickGel IFE Gels (10)	Fixative	1 vial

Acid Violet Stain (1 vial)	IgG	1 vial

Tris-Buffered Saline (1 pkg)	IgA	1 vial

Citric Acid Destain (1 pkg)	IgM	1 vial

Applicator Blade Assembly (20)	Kappa	1 vial

QuickGel Blotter C (20)	Lambda	1 vial

QuickGel IFE Blotter J (10)

QuickGel IFE Blotter Combs (10)

Materials provided but not contained in the kit:

Item		  Cat. No.

SPIFE 3000 Analyzer		1088

REP Prep		3100

SPIFE QuickGel IFE Cup Tray		3371

SPIFE IFE 9/15 Disposable Sample Cups		3363

SPIFE Gel Block Remover		1115

SPIFE QuickGel Rigid Antisera Template		3372

SPIFE QuickGel Electrodes		1111

SPIFE QuickGel Holder		3358

SPIFE QuickGel IFE Alignment Guide		3373

QuickGel IFE Templates and QuickGel Blotter A	552156

Materials and Supplies Needed but not Supplied:

Pipette and tips

10% Glacial acetic acid

0.85% saline

STEP BY STEP METHOD

I.	Sample Preparation 	
A.	Serum		
The patient serum samples are diluted 1:3 (1 part serum with 2 parts 0.85% saline) for serum protein lanes and diluted 1:5 
(1 part serum with 4 parts 0.85% saline) for immunofix lanes.	
However, due to desired sensitivity variations, serum samples may also be diluted as follows:

IgG = 1:5 to 1:10

IgA = undiluted to 1:5

IgM = undiluted to 1:5

k = 1:5 to 1:10

l = undiluted to 1:5

The more concentrated samples are more likely to prozone while the more diluted samples may not exhibit desired sensitivity.

B.	Urine		
Urine samples may be run diluted, neat or concentrated. Shake samples to homogenize. Centrifuge desired volume at 2000 x g for 5 minutes. Remove supernatant and concentrate as follows:

Total Protein (mg/dL)	Conc. Factor

< 50	100x

50-100	50x

100-300	25x

300-600	10x

> 600	5x

C.	Cerebrospinal Fluid		
Concentrate CSF to an IgG level of 100-200 mg/dL for typing oligoclonal bands in CSF. Use concentrated specimen for all patterns. CSF can only be applied to the gel by template method.

II.	Sample Application 	
A.	Blade Application for Serum and Urine

1.	Slide the Disposable Cup Strip into the appropriate Cup Tray (Cat. No. 3371).

2.	Pipette 17 mL of diluted serum or 20 mL urine into the shallow wells of the Cup Strip. Samples should be placed in the wells aligned with “SP, G, A, M, K, L”.

3.	Place the Cup Tray into the SPIFE 3000. Align the holes in the tray with the pins on the instrument.

4.	Remove two disposable Applicator Blade Assemblies from the packaging. Remove the protective guard from the blades by gently bending the protective piece back and forth until it breaks free.

5.	Place the Applicator Blades into the vertical slots in the Applicator Assembly numbered 6 and 12. 				
NOTE: The blade assembly will only fit into the slots in the Applicator Assembly one way; do not try to force the Applicator Blade into the slots.

6.	Place the SPIFE QuickGel IFE Alignment Guide on the chamber floor by aligning the round hole over the left pin and the obround hole over the right pin. The two small tabs must be bent upward.

7.	Carefully open one end of the pouch and remove one gel from the protective packaging. Reseal the pouch with tape to prevent drying of the gel. Remove the gel from the plastic mold and discard the mold.

8.	Dispense approximately 1 mL of REP Prep onto the left side of the electrophoresis chamber.

9.	Hold the gel so that the logo is turned to the left side of the chamber. Gently lay the gel down on the REP Prep, starting with the left side, aligning the notch in the gel backing so that it fits around the small tab of the Alignment Guide. Carefully align the right notch in the gel to fit the right upright tab.

10.	 Place a QuickGel Blotter C on the gel with the longer edge parallel with gel blocks. Gently blot the entire surface of the gel using slight fingertip pressure on the blotter, and remove the blotter.

11.	 Use lint-free tissue to wipe around the edges of the plastic gel backing, especially next to electrode posts, to remove excess REP Prep. Make sure no bubbles remain under the gel.

12.	Clean the QuickGel Electrodes with deionized water before and after each use. Wipe with a lint-free tissue.

13.	Place a QuickGel Electrode on the outside edge of each gel block inside the magnetic posts. Improper contact of the electrodes and the gel blocks may cause skewed patterns. Close the chamber lid. Proceed to Step III.

B.	Template Application for CSF

1.	Follow Steps II. A. 6-11.

2.	Remove one QuickGel IFE template from the package. Hold the template so that the small hole in the corner is toward the front right side of the SPIFE 3000.

3.	Carefully place the template on the gel aligning the template slits with the marks on each side of the gel backing. The center hole in the template should align with the indention in the center of the gel. If running two samples, place a second template on the gel.

4.	Apply slight fingertip pressure to the template making sure there are no bubbles under it. NOTE: If wearing rubber gloves to perform this step, place a QuickGel Blotter A over the template and then apply fingertip pressure to the template. Remove the blotter. Powder from the gloves can produce gel artifacts.

5. 	Clean the electrodes with deionized water before and after each use. Wipe with a lint free tissue.

6. 	Place a QuickGel electrode on the outside ledge of each gel block inside the magnetic posts. Improper contact of the electrodes and the gel blocks may cause skewed patterns. Close the chamber lid. Proceed to Step III.

III.	Electrophoresis	
If testing only serum samples, follow the instructions marked “Serum”. If testing only urine or serum with urine, follow instructions marked “Urine or Urine and Serum”. The blade used for serum application will be added after the second application of urine.

Due to variation in environmental conditions,
*a Blot 1 time range of 2 to 5 minutes is acceptable.
**an antisera absorption time of 1 to 3 minutes is acceptable.

Using the instructions provided in the Operator’s Manual, set up parameters as follows.

•	Serum

Electrophoresis Unit

1)	No prompt

Load sample 1	00:30	21°C	SPD6

2)	No prompt

Apply sample 1	00:30	21°C	SPD1	LOC 1

3)	No prompt

Electrophoresis 1	8:00	21°C	350V	60 mA

4)	Remove gel blks, apply antisera (continue)

Absorb 1	**2:00	21°C

5)	Remove excess antisera (continue)

Blot 1	*2:00	21°C

6)	Remove template, install blot (continue)

Blot 2	5:00	40°C

7)	Remove blotter, (continue)

Dry 1	8:00	50°C

8)	No prompt

END OF TEST

•	Urine or Urine and Serum

Electrophoresis Unit

NOTE: Serum and urine samples may be run on the same gel on different rows by pipetting 20 μL urine and 17 μL diluted serum into the cups. Change Step “5) No prompt” to “5) To Continue, (continue)”.

Place applicator blades into the slots that correspond to the urine sample. After the second urine application, the machine will beep and stop. Add an applicator blade into the remaining slot for serum samples. Press TEST SELECT/CONTINUE, the machine will apply and continue.

1)	No prompt

Load sample 1	00:25	21°C	SPD6

2)	No prompt

Apply sample 1	00:25	21°C	SPD6	LOC1	 			3)	No prompt

Load sample 2	00:25	21°C	SPD6

4)	No prompt

Apply sample 2	00:25	21°C	SPD6	LOC1

5)	No prompt

Load sample 3	00:25	21°C	SPD6

6)	No prompt

Apply sample 3	00:25	21°C	SPD6	LOC1

7)	No prompt

Absorb 1	 2:00	21°C

8)	No prompt

Electrophoresis 1	 8:00	21°C	350V	60 mA

9)	Remove gel blks, apply Antisera (continue)

Absorb 2	**2:00	21°C

10)	Remove excess Antisera (continue)

Blot 1	*2:00	21°C

11)	Remove template, install blot (continue)

Blot 2	5:00	40°C

12)	Remove blotter, (continue)

Dry 1	8:00	50°C

13)	No prompt

END OF TEST

1.	Press the TEST SELECT/CONTINUE buttons located on the Electrophoresis and Stainer sides of the instrument until the QG SERUM/URINE IFE option appears on the displays.

2.	With the appropriate display, press the START/STOP button. An option to either begin the test or skip the operation will be presented. Press START/STOP again to begin.

3.	The SPIFE 3000 will apply samples onto the gel and electrophorese, then beep when completed. Dispose of blades and cups as biohazardous waste. Proceed to Step IV.

•	CSF (Template Application)

Electrophoresis Unit

1)	Apply sample to template, (continue)

Absorb 1 	2:00 	21°C

2)	Blot and remove template, (continue)

Electrophoresis 1	8:00 	21°C 	350V	60 mA

3)	Remove gel blocks, apply antisera, (continue)

Absorb 2	**2:00 	21°C

4)	Remove excess antisera, (continue)

Blot 1	*2:00 	21°C

5)	Remove template, install blot (continue)

Blot 2	5:00 	40°C

6)	Remove blotter, (continue)

Dry 1	8:00 	50°C

7)	No Prompt

END OF TEST

1.	With QG SERUM/URINE IFE on the display, press the START/STOP button. An option to either begin the test or skip the operation will be presented. Press START/STOP to begin. Open the chamber lid.

2.	Place 3 mL of each sample onto the slits in the template (one protein and five immunofixation) for each patient. Apply the samples as quickly as possible.

3.	Close the chamber lid, and press the TEST SELECT/CONTINUE button for the electrophoresis chamber. Sample application will be timed for 2 minutes.

4.	After sample application is complete, open the chamber lid and gently blot the template with a QuickGel Blotter A.

5.	Carefully remove the blotter(s) and template(s) and discard as biohazardous waste.

6.	Close the chamber lid, and press the TEST SELECT/CONTINUE button to start electro­phoresis. SPIFE will beep when electrophoresis is complete.

•	Serum, CSF and Urine

Stainer Unit

1)	Plate Out, Holder In, Press (Continue)

Wash 1	00:30	REC=ON	Valve=1

2)	Plate In, Gel Holder In, Press (Continue)

Wash 2	10:00	REC=ON	Valve=1

3)	No prompt

Stain 1	4:00	REC=OFF	Valve=5

4)	No prompt

Destain 1	1:00	REC=ON	Valve=2

5)	No prompt

Destain 2	1:00	REC=ON	Valve=2

6)	No prompt

Dry 1	8:00	63°C

7)	No prompt

Destain 3	1:00	REC=ON	Valve=2

8)	No prompt

Dry 2	5:00	63°C

9)	No prompt

END OF TEST

IV.	Immunofixation

1.	When electrophoresis is complete, open the chamber lid. Remove the QuickGel electrodes.

2.	Using the SPIFE Gel Block Remover, remove and discard both gel blocks. Wipe around the edges of the gel to remove excess moisture.

3.	Holding the Antisera Template by the handles, gently place it onto the surface of the gel such that the round alignment hole is positioned on the pin to the left and the obround hole fits over the alignment pin on the right. No further pressure is needed.

4.	Quickly pipette 140 mL of Fixative and Antisera into the oval slots at the right end (anode) of each antisera channel in the template.

5.	Close the chamber lid and press the TEST SELECT/CONTINUE button to continue with antisera absorption. After the **2 minute absorption time, the SPIFE will beep.

6.	When antisera absorption is complete, open the chamber lid. Place one QuickGel Blotter Comb into the slots on the right end of the antisera channels such that the tips of the combs touch the gel. Close the chamber lid, and press the TEST SELECT/CONTINUE button. The preliminary blot will be timed for *2 minutes and the instrument will beep.

7.	Remove the Blotter Combs and the Antisera Template. Gently blot the gel with a QuickGel Blotter C and remove the blotter. Place a Blotter J on the surface of the gel. Place the Antisera Template on top of the Blotter J. Close the chamber lid and press the TEST SELECT/CONTINUE button. The final blot will be timed for 5 minutes.

8.	When the beep sounds, open the chamber lid and remove the Antisera Template and discard the blotter. Replace the QuickGel Electrodes on the ends of the gel to prevent curling. Close the chamber lid and press the TEST SELECT/CONTINUE button. The gel will be predried in the electrophoresis chamber.		
NOTE: Do not allow antisera to dry in the template. The Antisera Template should be cleaned with a mild biocidal detergent. The template may also be scrubbed with a soft brush to remove any antisera residue. Rinse with deionized water.

9.	After the gel has been predried, carefully remove the gel from the electrophoresis chamber.

V. Washing, Staining, and Destaining

1.	With the appropriate test name on the display, press the START/STOP button. An option to either begin the test or skip the operation will be presented. Press START/STOP to begin.

2.	Press the TEST SELECT/CONTINUE button. This will initiate the chamber prerinse cycle.

3.	After the chamber has been rinsed, remove the QuickGel Holder from the Stainer Unit. While holding the gel agarose side down, slide one side of the gel backing under one of the metal bars. Bend the gel backing so that the gel is bowed, and slip the other side under the other metal bar. The two small notches in the backing must fit over the small pins to secure the gel to the holder.

4.	Place the SPIFE QuickGel Holder with the attached gel facing backwards into the stainer chamber.

5.	Press the TEST SELECT/CONTINUE button to begin the staining process. The instrument will wash, stain, destain and dry the gel.

6.	When the process is completed, the instrument will beep. Carefully remove the SPIFE QuickGel Holder from the stainer because the metal piece on the holder will be hot.

PROCEDURE FOR QuickGel® CHAMBER

The following instructions are for using the QuickGel Chamber (Cat. No. 1284) for electrophoresis.

Materials Provided: The following materials needed for the procedure are contained in the QuickGel IFE Kits (Cat. No. 3551 or 3551T). Individual items are not available.

Cat. No. 3551T

QuickGel IFE Gels (10)	Fixative	1 vial

Acid Violet Stain (1 vial)	IgG	1 vial

Tris-Buffered Saline (1 pkg)	IgA	1 vial

Citric Acid Destain (1 pkg)	IgM	1 vial

QuickGel IFE Templates (20)	Kappa	1 vial

QuickGel Blotter A (20)	Lambda	1 vial

QuickGel Blotter C (20)

QuickGel IFE Blotter J (10)

QuickGel IFE Blotter Combs (10)

Cat. No. 3551

QuickGel IFE Gels (10)	Fixative	1 vial

Acid Violet Stain (1 vial)	IgG	1 vial

Tris-Buffered Saline (1 pkg)	IgA	1 vial

Citric Acid Destain (1 pkg)	IgM	1 vial

QuickGel IFE Applicator Blades (20)	Kappa	1 vial

QuickGel Blotter C (20)	Lambda	1 vial

QuickGel IFE Blotter J (10)

QuickGel IFE Blotter Combs (10)

QuickGel Dispo Sample Cups (10)

Materials provided but not contained in the kit:

Item		Cat. No.

QuickGel Chamber		1284

QuickGel IFE Rigid Antisera Template		3552

REP Prep		3100

QuickGel Gel Block Remover		1262

Staining Dish		4061

QuickGel Applicator		1265

QuickGel Applicator Base		1266

QuickGel Applicator Weights		1267

QuickGel Dispo Cup Tray		1268

QuickGel IFE Dispo Sample Cups (50/pkg)		1258

QuickGel Applicator Kit (includes Applicator,

Applicator Base, Weights and Cup Tray)		1274

QuickGel IFE Applicator Blades (10/pkg)		1272

Materials needed but not provided:

Pipette and tips

10% acetic acid

0.85% saline

Power Supply capable of providing at least 350 Volts

STEP BY STEP METHOD

I.	Sample Preparation	
Sample application for serum and urine may be done with applicator blades or templates. CSF must be done with templates only		
NOTE: Helena’s recommendation for sample dilutions are given below. However, due to desired sensitivity variations or sample volumes, serum samples may also be diluted more or less as needed. The more concentrated samples are more likely to prozone while the more diluted samples may not exhibit desired sensitivity.	
A.	Serum

1. Template Application		
The patient serum samples are diluted 1:5 (1 part serum with 4 parts 0.85% saline) for serum protein lanes and diluted 1:10 (1 part serum with 9 parts 0.85% saline) for immunofix lanes.

2. Blade Application		
The patient serum samples are diluted 1:3 (1 part serum with 2 parts 0.85% saline) for serum protein lanes and diluted 1:5 (1 part serum with 4 parts 0.85% saline) for immunofix lanes.

However due to desired sensitivity variations, serum samples may also be diluted as follows:

IgG	= 1:5 to 1:10

IgA	= undiluted to 1:5

IgM	= undiluted to 1:5

K	= 1:5 to 1:10

L	= undiluted to 1:5

B.	Urine (Template or Blade Application)		
Urine samples may be run diluted, neat or concentrated. Shake samples to homogenize. Centrifuge desired volume at 2000 x g for 5 minutes. Remove supernatant and concentrate as follows:

Total Protein (mg/dL)	Conc. Factor

< 50	100x

50-100	50x

100-300	25x

300-600	10x

> 600	5x

C.	Cerebrospinal Fluid (Template Method)		
Concentrate CSF to an IgG level of 100-200 mg/dL for typing oligoclonal bands in CSF. Use concentrated specimen for all patterns.

II.	Chamber Preparation

1.	The QuickGel Chamber must be plugged into a power supply. Set a timer for *7:30 minutes and the power at 350 Volts.

*An electrophoresis time of 7:00 to 8:00 minutes is acceptable.

2.	Snap the Electrophoresis Lid into place on the chamber.

3.	Ensure that the chamber floor is cool (room temperature) before starting the test.

III.	Gel Preparation

1.	Carefully cut open one end of the gel pouch. Remove one gel from the protective packaging. Fold and tape the open end of the pouch to prevent drying of the remaining gel. Remove the gel to be used from the plastic mold and discard the mold.

2.	Dispense approximately 1 mL of REP Prep onto the left side of the electrophoresis chamber.

3.	Hold the gel so that the logo is turned to the left side of the chamber. Place the left notch of the gel so that it fits the left pin of the chamber floor and gently roll the gel to the right side fitting the right notch to the right pin of the chamber floor. Use a lint free tissue to wipe around the edges of the gel backing to remove any excess REP Prep. Make sure that no bubbles remain under the gel.

4.	Place a QuickGel Blotter C on the gel with the longer sides parallel with the gel blocks. Gently blot the entire gel using slight fingertip pressure on the blotter. Remove the blotter.

IV.	Sample Application

A.	Template Application

1.	Remove two QuickGel IFE Templates from the package. Hold the template so that the small hole in the corner is toward the front right side of the chamber.

2.	Carefully place the template on the gel aligning the template slits with the marks on each side of the gel backing. The center hole in the template should align with the indention in the center of the gel. Two samples can be run on one gel.

3.	Apply slight fingertip pressure to the template making sure there are no bubbles under it. NOTE: If wearing rubber gloves to perform this step, place a QuickGel Blotter A over the template and then apply fingertip pressure to the template. Remove the blotter. Powder from the gloves can produce gel artifacts.

4.	Apply 3 mL of the appropriate serum dilution, CSF or urine to the template slits. Wait 2 minutes after the last sample application to allow proper absorption.

5.	Use the QuickGel Blotter A to gently blot the excess sample from the template. Then carefully remove the template.

6.	Close the lid, press the power switch to turn on the chamber and start the power supply.

B.	Blade Application

1.	Remove one QuickGel Applicator Blade from the packaging. If testing two samples, remove two Applicator Blades from the packaging.

2.	Urine samples must be applied on the gel three times. When testing serum with urine samples, serum application is made after the second urine application. Therefore the blade for serum application is not added until after the second urine application.

3.	Place an Applicator Weight on top of each Applicator Blade. If using two Applicator Blades, place them into the vertical slots A and C of the Applicator. One blade should be placed into the slot corresponding to cup placement.

4.	Place the appropriate number of QuickGel IFE Dispo Sample Cups into either or both Row A and Row C of the Dispo Cup Tray. Pipette 17 mL of dilutent serum or 20 mL urine into the Sample Cups and cover tray until ready to use. When ready, place the Dispo Cup Tray into the Applicator Base.

5.	If testing only serum samples, follow steps 6 through 11. If testing serum and urine on the same gel, perform steps 6 through 10 twice for the urine samples. Place the Modified Blade for serum samples into the applicator, and repeat steps 6 through 10 for a total of 3 applications for urine samples and 1 application for serum samples.

6.	While holding the white Applicator knob up, place the Applicator into the designated slot on the Applicator Base aligning the small red dots on the Applicator with those on the Base.

7.	Slowly lower the Applicator Knob allowing the blades to enter the sample cups, and immediately start a timer for 30 seconds.

8.	After 30 seconds, lift the Applicator Knob. Immediately place the Applicator into the slot on the chamber floor, aligning the red dots.

9.	Slowly lower the Applicator Knob to apply sample to the gel. Set a timer for 30 seconds.

10.	After the 30 seconds application, raise the Applicator Knob and remove the Applicator from the chamber.

11.	Close the lid, press the power switch to turn on the chamber and start the power supply.

V.	Electrophoresis and Staining

1.	Electrophorese the gel for *7:30 minutes at 350 Volts.

2.	Turn off the Power Supply and the QuickGel Chamber. Remove the lid.

3.	Using the QuickGel Gel Block Remover, carefully remove just the two gel blocks. Antisera leakage can occur if more than just the gel blocks are removed. Again, use a lint free tissue to wipe around the edges of the gel backing to remove any excess moisture.

4.	Replace the Electrophoresis Lid with the Drying Lid. Clean the two electrodes on the Electrophoresis Lid with deionized water after each use. Wipe with a lint free tissue.

5.	Holding the Antisera Template in the up position, place the template into the appropriate slots in the chamber.

6.	Gently lower the Antisera Template onto the surface of the gel.  No further pressure is needed.

7.	Quickly pipette 140 mL of Fixative and Antisera into the slots at the right end (anode) of each Antisera channel in the template.

8.	After 2 minutes (1 to 3 minutes is acceptable) absorption time, place one Blotter Comb into the same slots as the Antisera application. After 2 minutes, remove the blotter comb and the Antisera Template.

9.	Gently blot the gel with a QuickGel Blotter C and remove the blotter. Place a QuickGel Blotter J on the surface of the gel. Place the Antisera template on top of the Blotter J. After 5 minutes, remove the Antisera Template and Blotter J. Close the lid.

10.	Dry the gel for 8 minutes or until dry by turning on only the QuickGel Chamber. After drying is complete, turn the chamber off and remove the gel.

11.	Fill a Staining Dish with prepared stain. Fill another one with Destain solution. Fill a third dish with Tris-Buffered Saline (TBS).

12.	Place the gel in TBS and shake gently on a rotator for 10 minutes. Remove the gel from the TBS and allow it to drain on a blotter.

13.	Place the gel into the Staining Dish containing the prepared stain for 4 minutes. Remove the gel from the stain and allow it to drain on a blotter.

14.	Destain the gel in two consecutive washes of Destain solution. Use a gentle alternately rocking and swirling technique. Allow the gel to remain in each wash for 1 minute. The gel background should be completely clear. Tap the gel to remove the excess destain solution.

15.	Ensure that the chamber floor is clean. Replace the gel onto the QuickGel Chamber floor. Close the Drying Lid, turn on the QuickGel Chamber and dry for 10 minutes or until dry.

16.	Remove the gel from the chamber, and place it into the Destain again for 1 minute. Tap the gel to remove excess destain solution.

17.	Place the gel back into the chamber and dry for 5 minutes.

18.	Turn off the chamber and remove the gel.

Stability of the End Product: The completed, stained and dried immunofixation gel is stable for an indefinite period of time.

Quality Control: IFE controls may be required by federal, state and local regulations.

INTERPRETATION OF RESULTS

The majority of monoclonal proteins migrate in the cathodic (gamma) region of the protein pattern. But, due to their abnormality, they may migrate anywhere within the globulin region on protein electrophoresis. The monoclonal protein band on the immunofixation pattern will occupy the same migration position and shape as the monoclonal band on the reference protein electrophoresis pattern. The abnormal protein is identified by the corresponding antiserum used. When low concentrations of M-protein are present, the immunofixation band may appear on the stained background of the polyclonal immunoglobulin. A stained background may also appear when the M-protein is present along with a large polyclonal increase.

For an in-depth discussion of IFE interpretation, call Helena Laboratories toll free and request the free publication “Immuno- Fixation for the Identification of Monoclonal Gammopathies” Form R5.

LIMITATIONS

1. Antigen excess will occur if there is not a slight antibody excess or antigen/antibody equivalency at the site of precipitation. Antigen excess in IFE is usually due to a very high level of immunoglobulin in the patient sample. The dissolution of immunoprecipitation is manifested by a loss of protein at the point of highest antigen concentration, resulting in staining in the margins and leaving the central area with little demonstrable protein stain. In this case it may be necessary to adjust the protein content of the sample by dilution. Electrophoresing excessive amounts of antigen decreases resolution and requires higher concentrations of antibody. For optimum separation and sufficient intensity for visual detection, care must be taken in adjusting antibody content, sample concentration, time and voltage. The QuickGel ImmunoFix method has been optimally developed to minimize the antigen excess phenomenon.

2. Monoclonal proteins may occasionally adhere to the gel matrix, especially IgM. These bands will appear in all five antisera reaction areas of the gel. However, where the band reacts with the specific antisera for its heavy chain and light chain, there will be a marked increase in size and staining activity, allowing the band to be identified.

Further Testing Required:

Specimens containing a band on serum protein electrophoresis suggestive of a monoclonal protein, but which do not react with IgG, IgA or IgM antisera, may require further testing as follows:

1.	Serum samples which have a precipitin band with Kappa or Lambda Light Chain Antisera but none corresponding with IgG, IgA or IgM antisera may have a free light chain or they may have an IgD or IgE monoclonal protein. Such sera should be tested with ImmunoFix IgD and IgE antisera.

2.	A CRP band may be detected in patients with acute inflammatory response.7, 8 CRP appears as a narrow band on the most cathodic end of the high resolution agarose protein electrophoresis pattern. Evaluated alpha1 antitrypsin and haptoglobin (acute phase proteins) are supportive evidence for the presence of a CRP band. Patients with a CRP band will have a positive CRP by latex agglutination or an elevated quantitative CRP.

3.	Cerebrospinal fluid may contain a non-immunoglobulin band, referred to as gamma-trace, which migrates in the gamma region. Because gamma-trace is non-immunoglobulin in nature, it will not react with antisera against human immunoglbulins. Gamma-trace is often detected in normal cerebrospinal fluid. 9,10

PERFORMANCE CHARACTERISTICS

Serum, urine and CSF specimens containing monoclonal and polyclonal proteins were tested using the SPIFE ImmunoFix and the QuickGel ImmunoFix procedures. The test results showed good agreement between the methods.

BIBLIOGRAPHY

1.	Alfonso, E., Quantitation Immunoelectrophoresis of Serum Proteins, Clin Chem Acta, 10:114-122, 1964.

2.	Alper, C.A. and Johnson, A.M., Immunofixation Electrophoresis:A Technique for the Study of Protein Polymorphism. Vo Sang17:445-452, 1969.

3.	Alper, C.A., Genetic Polymorphism of Complement Componentsas a Probe of Structure and Function. Progress in Immunology, First International Congress of Immunology. Edited by New York, Academic Press, 609-624, 1971.

4.	Johnson, A.M. Genetic Typing of Alpha(1)-Antitrypsin in Immuno-fixation Electrophoresis. Identification of Subtypes of P.M., J Lab Clin Med, 87:152-163, 1976.

5.	Cawley, L.P. et al., Immunofixation Electrophoretic Technique Applied to Identification of Proteins in Serum and Cerebrospinal Fluid, Clin Chem, 22:1262-1268, 1976.

6.	Ritchie, R.F. and Smith, R., Immunofixation III, Application to the Study of Monoclonal Proteins, Clin Chem, 22:1982-1985,1976.

7.	Jeppsson, J.E., et al., Agarose Gel Electrophoresis, Clin Chem, 25(4):629-638, 1979.

8.	Killingsworth, L.M., et al., Protein Analysis, Diagnostic Medicine, 3-15, Jan/Feb., 1980.

9.	Keshgegian, A.A., et al., Oligoclonal Immunoglobulins in Cerebrospinal Fluid Multiple Sclerosis, Clin Chem, 26(9):1340-1345, 1980.

10.	Papadopoulos, N.M., et al., A Unique Protein in Normal Human Cerebrospinal Fluid, Clin Chem, 29(10):1842-1844, 1983.

For Sales, Technical and Order Information and Service Assistance, call 800-231-5663 toll free.

Helena Laboratories warrants its products to meet our published specifications and to be free from defects in materials and workmanship. Helena’s liability under this contract or otherwise shall be limited to replacement or refund of any amount not to exceed the purchase price attributable to the goods as to which such claim is made. These alternatives shall be buyer’s exclusive remedies.

In no case will Helena Laboratories be liable for consequential damages even if Helena has been advised as to the possibility of such damages.

The foregoing warranties are in lieu of all warranties expressed or implied including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose.

© 2023 Helena Laboratories, Corp.

Pro. 123

7/23(5)