PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

- It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

- The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

- The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

- PRODUCT {Test} NAME

- INTENDED USE and TEST TYPE (qualitative or qualitative)

- SUMMARY AND EXPLANATION

- PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

- REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

- INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

- SAMPLE COLLECTION/HANDLING

- PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

Actalyke® Whole Blood QC Kit

Cat. No. AQC–L	 For In Vitro Diagnostic Use	Store at 2-8°C.

Ranges / Bereiche / Intervalo de valores / Plages / Range / Amplitudes / Omfång

Tube Type	Level and Range (seconds)	Level and Range (seconds)

AQC-L	G-ACT	Level I  -	Level II -

Kit Lot No. 	Exp.

Vial Lot No. 	Level I: 	Level II:

H2O: 	CaCl2:

G-ACT ranges have been determined by assaying QC material on the Actalyke, Actalyke MINI and Actalyke XL Systems only.

Intended Use

The Actalyke Whole Blood QC Kit is a whole blood control product intended for performing routine quality control testing of the Actalyke Activated Clotting Time test system. Two product configurations are available to meet different customer needs and to simulate different clinically significant ACT test results. Cat. No. AQC-L contains Level I and Level II controls and is recommended to perform QC testing representative of normal and moderately heparinized patient specimens (such as those undergoing ECMO procedures or renal dialysis). Cat. No. AQC-H contains Level I and Level III controls and is recommended to perform QC testing representative of normal and highly heparinized patient specimens (such as those undergoing cardiovascular procedures).

Summary and Explanation

Actalyke Whole Blood QC kits are recommended for the verification of proper test system performance combined with operator technique. In the US according to the CLIA ‘88 regulations, the frequency of quality control testing is mandated. Since the ACT is categorized as moderately complex, ACT users must (1) perform two levels of quality control during each shift in which the test system is used clinically, and (2) subscribe to a Proficiency Testing program. To help users comply with these regulations, Actalyke Whole Blood QC Kits are available.

Actalyke Whole Blood QC kits are manufactured using non-human plasma and provide results that are similar to those expected clinically, however some differences in values will be observed as compared to whole blood patient samples.

An expected range of QC values, appropriate to specific Actalyke test tubes, is published for each lot of Actalyke Whole Blood QC kits. This expected range facilitates the analysis of QC results obtained at a given institution, and enables the user to determine if his/her QC results are falling within “acceptable” limits. In addition, a “QC Management Chart” is provided in each Actalyke Whole Blood QC Kit to help users track their quality control test results.

Storage and Stability

When stored at 2-8°C, all unopened vials are stable until the marked expiration date. Unreconstituted whole blood control vials are stable for seven days when stored at room temperature. After reconstitution, the whole blood control vials are stable for four hours at room temperature.

Reconstitution Water vials and Calcium Chloride vials are stable until the marked expiration date and may be stored at room temperature. To avoid contamination, use a clean needle and syringe for each test. If the reconstitution water or calcium chloride appears cloudy in the vial, there is evidence of contamination and it should be discarded.

Procedure

Materials Supplied

Cat. No. AQC-L (for use with G-ACT tubes)

• 1 Vial of Actalyke Whole Blood QC - Level I - Normal whole blood coagulation control, 3 mL lyophilized, containing a preparation of citrated animal plasma, hemoglobin, stabilizers and buffers.

• 1 Vial of Actalyke Whole Blood QC - Level II - Moderately elevated whole blood coagulation control, 3 mL lyophilized, containing a preparation of citrated animal plasma, hemoglobin, stabilizers and buffers.

• 1 Vial Reconstitution Water - Distilled water, 7.0 mL each

• 1 Vial Calcium Chloride - 0.02 M calcium chloride, 7.0 mL each

Contains no human material.

Materials required, not supplied

• Actalyke, Actalyke MINI, or Actalyke XL instrument

• Actalyke ACT test tubes (G-ACT tubes)

• Plastic syringes (3.0 mL & 1.0 mL) with needle

Step-By-Step

Note: If kept refrigerated, allow all Actalyke Whole Blood control vials, Reconstitution Water and Calcium Chloride vials to reach room temperature prior to use. This may take up to one hour.

1. Remove the seal on each vial.

2. Use a needle and syringe to puncture the rubber stopper of the Reconstitution Water vial. Aspirate 3.0 mL of Reconstitution Water and add it to the appropriate Actalyke Whole Blood control vial by puncturing the rubber stopper. Swirl gently to mix.

3. Allow the reconstituted Actalyke Whole Blood control vial to stand for 5 minutes swirling it gently at periodic intervals to assure complete reconstitution. At the end of 5 minutes, check material for complete dissolution. Allow to stand longer if not dissolved.

4. Use a needle and syringe to puncture the rubber stopper of the Calcium Chloride vial. Aspirate 0.25 mL of Calcium Chloride, and add it to the G-ACT test tube. Mix well to assure complete dispersion of the activator

5. Use a needle and syringe to aspirate 0.25 mL of the reconstituted Whole Blood control from the vial, and transfer it to the G-ACT test tube containing Calcium Chloride

6. Test the control in the same manner as the patient sample. Refer to the appropriate procedural insert for detailed instructions.

Limitations of the Procedure

Failure to obtain results that fall within the expected range of QC values may indicate product deterioration. Other factors that can also affect QC test results include technique and the temperature of the environment and materials used. If a QC test result falls outside the expected range, the test should be repeated using careful technique. If not, a study of each component of the system (i.e., Actalyke Whole Blood QC kit, Actalyke ACT test tube, and Actalyke instrument) should be performed so that the specific problem can be identified and corrected.

Performance Characteristics

In-house Precision Study

Level I

Site		day 1	day 2	day 3

1	rep 1	160	131	147
	rep 2	153	142	151

2	rep 1	152	159	153
	rep 2	141	128	155

3	rep 1	180	160	156
	rep 2	184	158	146

mean: 153.11	SD: 13.99	%CV: 9.14

Level II

Site		day 1	day 2	day 3

rep 1	196	221	204
	rep 2	204	218	197

2	rep 1	202	194	208
	rep 2	199	190	202

3	rep 1	203	234	218
	rep 2	208	236	216

mean: 208.33	SD: 13.03	%CV: 6.25

Performance Characteristics

In-house Precision

Level

Site

day

rep (replicate)

mean

SD (Standard Deviation)

%CV (percentage coefficient of variation)

Bibliography

1. NCCLS AST2-P, March 1995, Point-of-Care In Vitro Diagnostics (IVD) Testing; Proposed Guideline.

2. Statland B and Westgard, J: Quality control: Theory and Practice, Clinical Diagnosis and Management by Laboratory Methods, ed, Henry, J (Philadelphia: WB Saunders Co., 1984) pp. 74-93.

3. Department of Health & Human Services, Center for Laboratories, HSQB, Yost Letter Feb., 15 1996: Electronic Quality Control (QC) Policy; CLIA requirements: 493.1202(e)(4).

4. Colton C: Statistics in medicine (Boston: Little, Brown and Co. 1974) pp. 289-313.

Manufacturer

European authorized representative
Helena Laboratories UK, Ltd.
Colima Avenue
Sunderland Enterprise Park
Sunderland SR5 3XB
England

Form 377

1/04(5)