PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

- It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

- The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

- The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

- PRODUCT {Test} NAME

- INTENDED USE and TEST TYPE (qualitative or qualitative)

- SUMMARY AND EXPLANATION

- PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

- REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

- INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

- SAMPLE COLLECTION/HANDLING

- PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

V8 Nexus Hemoglobin UltraScreen Procedure

Cat. No. 1828

intended use

The V8 Nexus Hemoglobin UltraScreen method is designed for the separation of normal hemoglobins (A, A2, and F) in human blood samples, and for the detection of major hemoglobins variants (S and C) by using a capillary zone electrophoresis (CZE) buffer with the V8 instrument. The V8 Nexus Hemoglobin UltraScreen test is indicated for use in patients two (2) years of age and older. This test is designed for in-vitro diagnostic use only in conjunction with other laboratory and clinical findings.

The V8 instrument is an automated analyzer which performs a complete hemoglobin profile for quantitative analysis of the normal hemoglobin fractions A, A2 and F and for the detection of major hemoglobin variants S and C. The assay is performed on the hemolysate of venous whole blood collected in tubes containing K2EDTA as the anticoagulant.

For In Vitro Diagnostic use.

Rx Only

summary

Over 400 variant hemoglobins are now known, some of which may cause serious clinical effects, especially when present in the homozygous state, or in combination with another hemoglobin type. Wintrobe1 divides the abnormalities of hemoglobin synthesis into 3 groups:

(1) Production of an abnormal protein molecule.

(2) Reduction in the amounts of normal protein synthesis.

(3) Developmental anomalies.

PRINCIPLE

The V8 instrument is a counter top electrophoresis system capable of separating variant and normal hemoglobins. Capillary electrophoresis is achieved using a microbore, fused silica capillary filled with an appropriate electrolyte medium under high voltage. Separation of the charged analytes is due to differential migration in an electrical field, whereby the charged particles migrate toward an electrode with an opposite charge. The application of the EOF (electro-osmotic flow) and Coulomb force allows all charged species to be detected at one end of the capillary within zones of migration, due to mass to charge ratio separation alongside EOF induced movement towards the optical detector.

With so many variant hemoglobins, there is no one technique which can unequivocally identify each hemoglobin variant. The capillary zone electrophoresis method available on the V8 system will separate the more common variants from the normal A peak. The run positions of variant hemoglobins can be used to give an initial identification of the variant. For a more positive identification another technique or protein sequencing must be used as per IFCC guidance.

REAGENTS

1.	V8 Hemoglobin UltraScreen Buffer (1 x 500 mL)	
Ingredients: The buffer contains Taurine, L-Arginine, Tris, and Proclin 950.	
WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. DO NOT INGEST.		
Preparation for Use: The reagent is ready for use as packaged.	
Storage and Stability: The buffer should be stored at 15-30°C and is stable until the expiration date indicated on the label. Opened bottles placed on the V8 instrument are stable for 5 months.  DO NOT FREEZE THE LIQUID.	
Signs of Deterioration: The buffer should be a clear, colorless solution. Discard if solution shows extreme color change or cloudiness indicating microbial contamination.

2.	V8 Hemoglobin UltraScreen Diluent (2 x 20 mL)	
Ingredients: The diluent contains Triton X-100 and Proclin 950.	
WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. DO NOT INGEST.

Preparation for Use: The reagent is ready for use as packaged.	
Storage and Stability: The diluent should be stored at 15-30°C and is stable until the expiration date indicated on the label. Opened bottles placed on the V8 instrument are stable for 5 months. DO NOT FREEZE THE LIQUID.

Signs of Deterioration: The diluent should be a clear, colorless solution. Discard if solution shows extreme color change or cloudiness indicating microbial contamination.

INSTRUMENT

The Helena V8 Nexus capillary electrophoresis analyzer must be used for sample analysis. Refer to the Operator’s Manual for detailed instructions.

SPECIMEN COLLECTION AND HANDLING

Specimen: Freshly collected K2EDTA anticoagulated whole blood is the specimen of choice.

Specimen Storage and Stability: Samples can be stored at 2-8ºC for up to 7 days. However, best results are obtained with fresh samples as the formation of degradation product will complicate interpretation of the resulting hemoglobin trace.

Interfering Factors

Inaccurate results may be obtained on specimens left uncovered, due to evaporation. Samples older than 7 days cannot be used. Aged samples containing E variants can sometimes appear with a degradation peak to the right of the A2 peak. This degradation peak may be marked as A2 and therefore cause other peaks to be mislabeled in the trace. This can be corrected by manually locking peaks through the "edit peaks" right click menu. HbE is a hemoglobin variant with a similar electrophoretic mobility to HbA2. As a result of this, HbE increases the A2% in both heterozygous and homozygous expression. This interference is well documented and is present in HPLC and CZE methods. Interference should be taken into account when assessing a reference range for compound expression of beta thalassemia and HbE. As a result, a beta thalassemia cut off in the presence of HbE should be established and validated in the local population. HbA2 may be elevated in the presence of HbS or HbC, Methemoglobin migrates in the position of HbF in older samples.

Common interfering factors including lipids and bilirubin were tested using CLSI EP07-A2 guidelines.2 Lipid concentrations approaching 25 g/L and bilirubin concentrations approaching 
25 mg/dL demonstrated no qualitative or quantitative interference on the performance of the UltraScreen assay.

USAGE MATRIX

Approximate yield per bottle according to number of tests is given below:

40+ tests per session:		300 tests per bottle

32 tests per session:		272 tests per bottle

24 tests per session:		227 tests per bottle

PROCEDURE

Materials provided: The following materials needed for the procedure are contained in the V8 Hemoglobin UltraScreen Kit (Cat. No. 1828). Individual items are not available.

V8 Hemoglobin UltraScreen Buffer (1 x 500mL)

V8 Hemoglobin UltraScreen Diluent (2 x 20mL)

V8 Filter Unit (1)

V8 Sample Cups (3 pkg)

Materials provided but not contained in the above kit:

Items	Cat. No.

V8 Nexus CE Analyzer 	1825

V8 Storage Buffer	1831

V8 Maintenance Buffer	1832

V8 Clinical Waste Drawer Inserts	1820

V8 AFSA2 Hemo Control	1812

STEP-BY-STEP METHOD

For correct installation of all consumables, please refer to the V8 Operator Manual or Platinum Touchscreen Operator Manual.

1.	Before switching on the V8, ensure that Storage Buffer, Maintenance Buffer, disposable cups and waste container are onboard and in their correct positions.

2.	Switch on the V8 instrument, launch Platinum, and begin a new V8 session.  In Platinum, ensure that Hb UltraScreen is selected as the default method or use test ordering to assign the Hb UltraScreen assay to specific samples.

3.	If prompted by the V8 instrument, install the V8 Hemoglobin Ultrascreen Buffer bottle into the fluid bottle compartment and verify that it is in the correct position.		
NOTE:  For correct installation, apply a fresh filter unit to the inlet pipe of the buffer bottle connector before installing it on the V8 instrument.

4.	Installation of the V8 Hemoglobin UltraScreen Buffer bottle will automatically open a prompt in Platinum. Scan or enter the barcode number into the buffer position field matching the bottle position.

5.	Scan or enter the barcode number for the V8 Hemoglobin UltraScreen Diluent into Platinum, and verify that the location of the reagent information corresponds with the intended vial location onboard the V8 instrument.

6.	Remove the lid from the V8 Hemoglobin UltraScreen Diluent and ensure that the bottle is placed correctly in the reagent bottle area.

7.	When the V8 instrument is ready to accept samples for operation, the instrument will pulse red.

8.	Invert primary collection tubes several times to mix the whole blood.  Remove caps and load the sample tubes into the sample racks, ensuring that the barcodes are visible through the rack window.

9.	Load sample racks onto the left-hand side of the V8 sample transport area and close the lid.

NOTE: The analyzer should not have more than 2 racks onboard at a time due to the settling of red blood cells.

10.	The V8 instrument will automatically commence analysis of all loaded samples, and the results will be transferred to Platinum.

11.	After analysis, and if required, initiate V8 shut-down mode by switching off the instrument.

NOTE: It is important that the V8 is post-conditioned correctly at the end of the day. If the buffer port is to be left empty after V8 Hemoglobin UltraScreen Buffer usage, it is recommended to purge with V8 Storage Buffer before post-condition to prevent the line from drying out.

QUALITY CONTROL

The V8 AFSA2 Hemo Control (Cat. No. 1812) is to be used as a quantitative and/or qualitative control for the Hemoglobin UltraScreen assay on the V8 capillary electrophoresis system. Refer to the package insert provided with the control for the appropriate assay values and migration patterns.

INTERPRETATION OF RESULTS

Most hemoglobin variants cause no discernible clinical symptoms, so are of interest primarily to research scientists. Variants are clinically important when their presence leads to sickling disorders, thalassemia syndromes, lifelong cyanosis, hemolytic anemias, erythrocytosis, or if the heterozygote is of sufficient prevalence to warrant genetic counseling. The combinations of HbSS, HbSD-Los Angeles, and HbSO-Arab lead to serious sickling disorders.3  Several variants including HbH, E-Fort Worth and Lepore cause a thalassemic blood picture.3

The two variant hemoglobins of greatest importance in terms of frequency and pathology are HbS and HbC.3 Sickle cell anemia (HbSS) first manifests itself at about 5-6 months of age. The clinical course presents agonizing episodes of pain and temperature elevations with anemia, listlessness, lethargy and infarct in virtually all organs of the body. The individual with homozygous HbCC suffers mild hemolytic anemia which is attributed to the precipitation or crystallization of HbC within the erythrocytes. Cases of HbSC disease are characterized by hemolytic anemia that is milder than sickle-cell anemia.

The thalassemias are a group of hemoglobin disorders characterized by hypochromia and microcytosis due to the diminished synthesis of one globin chain (α or β) while synthesis of the other chain proceeds normally.4,5 This unbalanced synthesis results in unstable globin chains. These precipitate within the red cell, forming inclusion bodies that shorten the life span of the cell. In α-thalassemias, the α-chains are diminished or absent, and in the β-thalassemia, the β-chains are affected. Another quantitative disorder of hemoglobin synthesis, hereditary persistent fetal hemoglobin (HPFH), represents a genetic failure of the mechanisms that turn off gamma chain synthesis at about four months after birth, which results in a continued high percentage of HbF. It is a more benign condition than the true thalassemias and persons homozygous for HPFH have normal development, are asymptomatic and have no anemia.5

The most common hemoglobin abnormalities:

Sickle Cell Trait

This is a heterozygous state showing HbA, HbS and a normal amount of HbA2. Results show hemoglobins in the HbA and HbS migratory positions (zones).

Sickle Cell Anemia

This is a homozygous state showing almost exclusively HbS, although a small amount of HbF may also be present.

Sickle-C Disease

This is a heterozygous state demonstrating HbS and HbC.

Sickle Cell-Thalassemia Disease

This condition shows HbA, HbF, HbS and HbA2.

In Sickle Cell β°-Thalassemia HbA is absent.

In Sickle Cell β+-Thalassemia HbA is present in reduced quantities.

Thalassemia-C Disease

This condition shows HbA, HbF and HbC.

C Disease

This is a homozygous state showing almost exclusively HbC.

Thalassemia Major

This condition shows HbA, HbF and HbA2.

LIMITATIONS

1.	The use of aged or frozen samples may produce additional artifact peaks caused by hemoglobin degradation products.

2.	Globin chain analysis (both acid and alkaline) and structural studies may be necessary in order to positively identify some of the rarer hemoglobin variants.

3.	It is recommended that the result is considered as part of the entire clinical picture.

4.	Due to the resolution limits of capillary zone electrophoresis, it is possible that some hemoglobin variants may co-migrate with HbA or other variants using this method.

5. The V8 Hemoglobin UltraScreen kit has not been evaluated for use with dried blood spots (DBS) and the neonate subpopulation pediatric group.

6. A second test method is necessary for confirmation of abnormal hemoglobins detected.

7. Elevated HbA2 results should be interpreted with caution in patients with abnormal hemoglobins.

REFERENCE values

A reference range was established from 132 healthy individuals from 3 independent institutions using CLSI C28-A3 measurements.6

HbA: 96.6-97.9%

HbA2: 2.1-3.4%

Note: the reference range for HbF can be below the reportable limit of the assay.

Helena Laboratories recommends each laboratory establish its own reference range.

PERFORMANCE CHARACTERISTICS

Precision/Reproducibility

Lot to Lot:

Precision studies were performed using one lot of AFSA2 control, three lots of V8 Hemoglobin UltraScreen reagent kits, one V8 Nexus instrument and one operator. The tests were run 5 days x 2 runs/day x 8 reps/run. Each hemoglobin fraction has 240 total determinations (80 per fraction/lot). Within-run, between-run, between-day, between-lot and total precision were calculated. Acceptance precision criteria are all CVs for all Hemoglobin fractions ≤ 10%.

All hemoglobin fractions met acceptance criteria demonstrating acceptable lot to lot precision.

Within-run	Between-run   Between-day       Between-lot	       Total

Fraction	N	Mean 	SD	CV	SD	CV	SD	CV	SD	CV	SD	CV

%

A	240	42.9	0.26	0.6	1.63	3.8	0.26	0.6	0.43	1.0	1.72	4.0

F	240	33.3	0.17	0.5	1.03	3.1	0.20	0.6	0.27	0.8	1.10	3.3

S	240	21.8	0.26	1.2	0.57	2.6	0.37	1.4	0.57	2.6	0.92	4.2

A2	240	2.0	0.07	3.5	0.07	3.4	0.09	4.4	0.09	4.5	0.14	7.0

Instrument to Instrument:

Precision studies were performed using fresh and refrigerated venous K2EDTA patient samples with known hemoglobin variants. Normal HbAA2, HbAF (elevated F), HbASA2 (elevated S and A2) and HbAA2C (elevated C and A2) were analyzed over the course of 5 days x 2 runs/day x 8 reps per run study. One V8 UltraScreen reagent kit lot, three V8 Nexus instruments and a single operator were employed. Each sample has 240 total determinations per hemoglobin fraction (80 per instrument). Within-run, between-run, between-day, between-instrument and total precision were calculated. Acceptance precision criteria for patient samples are as all CVs for all Hemoglobin fractions ≤ 10%.

All hemoglobin fractions met acceptance criteria demonstrating acceptable precision.

Within-run	Between-run   Between-day         Between-	           Total

Instrument

HbA	N	Mean 	SD	CV	SD	CV	SD	CV	SD	CV	SD	CV

Sample		%

1	240	96.9	0.0	0.0	0.0	0.0	0.10	0.1	0.10	0.1	0.10	0.1

2	240	34.8	0.35	1.0	0.66	1.9	0.42	1.2	0.42	1.2	0.87	2.5

3	240	61.0	0.37	0.6	0.12	0.2	0.43	0.7	0.67	1.1	0.79	1.3

4	240	58.0	0.35	0.6	0.29	0.5	0.58	1.0	0.70	1.2	0.93	1.6

Within-run	Between-run   Between-day         Between-	           Total

Instrument

HbF	N	Mean 	SD	CV	SD	CV	SD	CV	SD	CV	SD	CV

Sample		%

2	240	65.2	0.33	0.5	0.65	1.0	0.46	0.7	0.39	0.6	0.85	1.3

Within-run	Between-run   Between-day         Between-	           Total

Instrument

HbA2	N	Mean 	SD	CV	SD	CV	SD	CV	SD	CV	SD	CV

Sample		%

1	240	3.1	0.05	1.5	0.00	0.1	0.05	1.6	0.05	1.6	0.07	2.2

3	240	3.6	0.05	1.3	0.01	0.4	0.05	1.3	0.05	1.3	0.07	1.9

4	240	5.5	0.07	1.2	0.06	1.1	0.07	1.2	0.07	1.2	0.12	2.1

Within-run	Between-run   Between-day         Between-	           Total

Instrument

HbS	N	Mean 	SD	CV	SD	CV	SD	CV	SD	CV	SD	CV

Sample		%

3	240	34.5	0.35	1.0	0.17	0.5	0.38	1.1	0.59	1.7	0.76	2.2

Within-run	Between-run   Between-day         Between-	           Total

Instrument

HbC	N	Mean 	SD	CV	SD	CV	SD	CV	SD	CV	SD	CV

Sample		%

4	240	36.5	0.37	1.0	0.29	0.8	0.58	1.6	0.73	2.0	0.95	2.6

LINEARITY / SENSITIVITY

The linearity of the V8 Nexus UltraScreen procedure was evaluated following CLSI EP06-A guidelines.7 The studies were performed using the a single V8 Nexus Hemoglobin UltraScreen reagent kit, a single V8 Nexus and venous K2EDTA samples containing different levels of each hemoglobin fraction (9 levels for each fraction). The tests were determined to be linear over the following ranges: HbA: 3.7-97.2%, HbF: 1.1-68.7%, HbS: 5.8-78.8%, HbA2: 1.7-7.6%, HbC: 1.4-42.6%.

Correlation Studies

Clinical correlation studies were performed at three separate institutions. HbA, HbA2, HbF, HbS, and HbC fractions were analyzed with the V8 Nexus Hemoglobin UltraScreen method versus the predicate capillary electrophoresis method. The degree of agreement with the predicate is outlined below.

Fraction	N	Range	R	Slope	95% CI	Intercept	95% CI

HbA	320	6.8-97.2	0.999	0.990	0.984,0.998	0.78	0.12,1.40

HbA2	412	1.6-6.9	0.957	1.000	1.000,1.000	0.05	0.05,0.05

HbF	175	1.1-68.0	0.993	1.000	0.985,1.020	0.70	0.52,0.73

HbS	143	9.6-78.6	0.994	0.929	0.909,0.948	2.40	1.67,3.22

HbC	33	23.6-42.3	0.975	1.000	0.925,1.103	0.60	-2.70,3.06

BIBLIOGRAPHY

1. Wintrobe, Maxwell M., Clinical Hematology, 6th Edition, Lea and Febiger, Philadelphia, 1967.

2. CLSI. Interference Testing in Clinical Chemistry; Approved Guideline—Second Edition. CLSI document EP07-A2. Wayne, PA:  Clinical and Laboratory Standards Institute; 2005.

3. Fairbanks, V.F., Diagnostic Medicine, Nov/Dec., 53-58, 1980.

4. Weatherall, D.J. and Clegg, J.B., The Thalassemia Syndromes, Blackwell Scientific Publications, Oxford, 1972.

5. Lehman, H. and Huntsman, R.G., Man’s Haemoglobins, J.B. Lippincott Co., Philadelphia, 1974.

6. CLSI. Defining, Establishing, and Verifying Reference Intervals in the Clinical Laboratory; Approved Guideline –Third Edition. CLSI document C28-A3. Wayne, PA:  Clinical Laboratory Standards Institute; 2008.

7. CLSI. Evaluation of the Linearity of Quantitative Measurement Procedures:  A Statistical Approach; Approved Guideline. CLSI document EP06-A. Wayne, PA:  Clinical Laboratory Standards Institute; 2003.

For Sales, Technical and Order Information and Service Assistance, call 800-231-5663 toll free.

Helena Laboratories warrants its products to meet our published specifications and to be free from defects in materials and workmanship. Helena’s liability under this contract or otherwise shall be limited to replacement or refund of any amount not to exceed the purchase price attributable to the goods as to which such claim is made. These alternatives shall be buyer’s exclusive remedies.

In no case will Helena Laboratories be liable for consequential damages even if Helena has been advised as to the possibility of such damages.

The foregoing warranties are in lieu of all warranties expressed or implied including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose.

Pro 007

5/23(3)