PLEASE READ!!

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

In response to customer requests, Helena is pleased to provide the text for procedural package inserts in a digital format editable for your use.  The text for the procedure you requested begins on page three of this document.  Helena procedures contain the content outlined in the NCCLS (GP2-A#) format, except in the order sequence required by FDA regulations.  As the NCCLS format is a guideline, you may retain these procedures as developed by the manufacturer (adding your title/authorization page) or manipulate the text file to produce your own document, matching the NCCLS section order exactly, if preferred.

We also provide the procedure in an Adobe Acrobat PDF format for download at www.helena.com as a “MASTER” file copy.  Below are the specifications and requirements for using these digital files.  Following the specifications is the procedure major heading sequence as given in the FDA style.  Where there is a difference in order, or other notation in the outline, this will be indicated in braces { }.

WHAT YOU NEED TO KNOW:

1)	These files represent the most current revision level to date.  Your current	product inventory could contain a previous revision level of this procedure.

2)	The Microsoft Word document provides the text only from the master procedure, in a single-column format.

-	It may not contain any illustrations, graphics or captions that may be part of the master procedure included in the kit.

-	The master procedure may have contained special formatting characters, such as subscripts, superscripts, degree symbols, mean symbols and Greek characters such as alpha, beta, gamma, etc.  These symbols may or may not display properly on your desktop.

-	The master procedures may also contain columns of tabbed data.  Tab settings may or may not be displayed properly on your desktop.

3)	The Adobe Acrobat PDF file provides a snapshot of the master procedure in a printable 8.5 x 11” format.  It is provided to serve as a reference for accuracy.

4)	By downloading this procedure, your institution is assuming responsibility for modification and usage.

HELENA LABORATORIES

PROCEDURE DOWNLOAD END USER AGREEMENT

HELENA LABORATORIES LABELING – Style/Format Outline

1)	PRODUCT {Test} NAME

2)	INTENDED USE and TEST TYPE (qualitative or qualitative)

3)	SUMMARY AND EXPLANATION

4)	PRINCIPLES OF THE PROCEDURE

{NCCLS lists SAMPLE COLLECTION/HANDLING next}

5)	REAGENTS (name/concentration; warnings/precautions; preparation; storage; environment; Purification/treatment; indications of instability)

6)	INSTRUMENTS required – Refer to Operator Manual (... for equipment for; use or function; Installation; Principles of operation; performance; Operating Instructions; Calibration* {*is next in order for NCCLS – also listed in “PROCEDURE”}’ precautions/limitations/hazards; Service and maintenance information

7)	SAMPLE COLLECTION/HANDLING

8)	PROCEDURE

{NCCLS lists QUALITY CONTROL (QC) next}

9)  RESULTS (calculations, as applicable; etc.)

10)  LIMITATIONS/NOTES/INTERFERENCES

11)  EXPECTED VALUES

12)  PERFORMANCE CHARACTERISTCS

13)  BIBLIOGRAPHY (of pertinent references)

14)  NAME AND PLACE OF BUSINESS OF MANUFACTURER

15)  DATE OF ISSUANCE OF LABELING (instructions)

For Sales, Technical and Order Information, and Service Assistance, 
call Helena Laboratories toll free at 1-800-231-5663.

Form 364

Helena Laboratories

1/2006 (Rev 3)

Thromboplastin LI Prothrombin Time Reagent

INTENDED USE

Thromboplastin LI is intended for use in a one-stage prothrombin time (PT) test on citrated human plasma. The PT test is a quantitative assay used in the general patient population for routine screening to detect deficiencies in the extrinsic pathway of coagulation. The PT test is also used to monitor oral anticoagulant therapy and should be used in a clinical laboratory by qualified laboratory personnel.

SUMMARY AND PRINCIPLE

The one-stage prothrombin time test (PT) has become the basic coagulation screening test for the diagnosis of congenital and acquired deficiencies of Factors II, V, VII and X1, 2. Oral anticoagulant drugs inhibit hepatic synthesis of the vitamin K dependent clotting Factors II, VII, IX and X. Therefore, the prothrombin time test is appropriate to monitor oral anticoagulant therapy since it is sensitive to three of the four factors involved3, 4 . It can be modified to quantitate factor levels using appropriate factor deficient substrate plasmas. Tissue thromboplastin, in the presence of calcium ions and Factor VII, activates the extrinsic pathway of coagulation. When a mixture of tissue thromboplastin and calcium ions is added to normal anti-coagulated plasma, the clotting mechanism is initiated and a clot will form within a specified time period. If a deficiency exists within the extrinsic pathway, the time required for clot formation will be prolonged. The degree of prolongation is proportional to the severity of single factor deficiency, or in a cumulative deficiency of all the factors involved.

REAGENTS

1.	Thromboplastin LI Reagent Ingredients: The reagent contains a liquid saline extract of rabbit brain, calcium ions, preservatives and stabilizers.

WARNING: FOR IN-VITRO DIAGNOSTIC USE ONLY. Preparation for Use: Reagent comes ready for use. Invert vial

3-4 times before use to ensure homogeneity. Storage and Stability: Reagent should be stored at 2 to 8°C and is stable until the expiration date indicated on the vial. After opening, the original activity is stable for fourteen (14) days at 2 to 8°C. DO NOT FREEZE. Signs of Deterioration: The reagent is a fine suspension of rabbit brain particles. Large flaky particles in the suspension or prolonged prothrombin times on testing normal plasma or controls may indicate product deterioration.

INSTRUMENTS

Prothrombin time test may be done by accepted manual methods or by using mechanical or photo-optical coagulation instruments such as the Cascade M and Cascade M-4.

SPECIMEN COLLECTION AND HANDLING NOTE: After initial whole blood collection with 3.2% sodium citrate, all test tubes, syringes and pipettes used during testing should be plastic. Specimen: Plasma obtained from whole blood anti-coagulated with 0.1 M (3.2%) sodium citrate. Specimen Collection: Nine parts freshly collected whole blood should be immediately added to one part anticoagulant.
Specimen Preparation: Specimens may be stored uncentrifuged or centrifuged with plasma remaining on the cells in an unopened tube at 18 to 24°C (CLSI H21-4, 2003). To obtain a plasma sample, the capped specimen tube should be centrifuged at 1500 x g for 15 minutes and the plasma removed with a plastic pipette, placed in a plastic tube and capped. PT assays should be completed within 24 hours of collection. If testing is not completed within 24 hours, plasma should be removed from cells and frozen at -20°C for up to 2 weeks or -70°C for up to 6 months.

Storage and Stability: Before and during testing, the plasma sample should be maintained in the plastic tubes at 2 to 8°C to ensure stability of the factors. If testing is delayed for more than 2 hours, the plasma may be stored at -20°C for two weeks or at -70°C for up to one month. Frozen samples should be thawed rapidly at 37°C before testing.

PROCEDURE Materials Provided: Material needed for prothrombin time tests are provided in the following packaging configurations:

Cat. No.

Thromboplastin LI Reagent (10 x 5 mL vials) 	5248

Thromboplastin LI Reagent (10 x 10 mL vials) 	5249

Materials and Equipment Required, but not provided:

1. Coagulation Instrument or 37°C Water Bath and timer

2. Reaction Cups or plastic test tubes (12 x 75 mm)

3. Pipetting device to deliver 200 µL or 100 µL

4. Centrifuge

5. Distilled or deionized water

6. Control plasmas: Norm-Trol 1 – Cat. No. 5186 Ab-Trol 2 - Cat. No. 5187 Ab-Trol 3 - Cat. No. 5183

PROCEDURE NOTE: Throughout the procedure, all test tubes, syringes, and pipettes should be plastic.

Automated and Semi-Automated Methods

If using an instrument to perform this test, refer to the appropriate Instrument Operator’s Manual for detailed instructions.

Quality Control:

Each laboratory should establish a quality control program that includes normal and abnormal controls to evaluate instrument, reagent and technologist performance. The normal and abnormal controls should be tested daily prior to performing tests on patient plasmas. Monthly quality control charts (Levy Jennings) are recommended to determine the mean and standard deviation of each of the daily control plasmas. A normal control such as Norm-Trol 1, and abnormal level 1 and level 2 controls such as Ab-Trol 2 and Ab-Trol 3 are recommended. If the controls do not perform within their reference ranges, patient results should be considered invalid and not reported.

EXPECTED VALUES

A reference range study was conducted using triplicate specimens from 120 normal healthy adults. Approximately equal numbers of males and females were used. The PT results were as follows:

Mean (seconds)	Range + 2 SD

Photo-optical 	12.3 	9.7 - 14.9 seconds

Mechanical 	12.6 	11.1 - 14.1 seconds

These values should only serve as guidelines. Because differences may exist between instruments, laboratories, and local populations, it is recommended that each laboratory establish its own reference range of expected prothrombin time results.

DETERMINATION OF INR

The World Health Organization (WHO) recommends the use of the International Normalized Ratio (INR) instead of direct reporting of the PT times when monitoring patients undergoing Oral Anticoagulant Therapy. This allows patient results to be compared between different laboratories that may be using reagents with different sensitivities.

The INR is calculated by using the ratio of the Patient PT to the mean of the normal reference range raised to the power of the reagent International Sensitivity Index (ISI).

INR = (Patient PT/mean of normal range)ISI

The ISI is assigned by comparison to a highly sensitive WHO thromboplastin standard reference material. The lower the ISI for

a given reagent the more sensitive the reagent is to coagulation factors. The lot specific ISI value for Thromboplastin LI can be

found on the enclosed ISI calculation sheet.

RESULTS

The results of the Prothrombin Time tests should be reported to the nearest tenth of a second. Results greater than the upper limits of the range should be considered abnormal and follow-up testing should be performed. PT values below the lower limits of the range may indicate a compromised sample, and a new sample should be collected.

LIMITATIONS

Expected values for the prothrombin time test will vary from one laboratory to another, depending on several variables. These include the method of clot detection, temperature, pH, sample collection technique, type of anticoagulant and time and method of plasma storage. Therefore, laboratories should establish their own expected values for patients and well defined performance standards for control plasmas. The use of icteric, lipemic, or hemolyzed sample should be avoided due to possible interference especially when using photo-optical instruments. The impact of other therapeutic drugs, in addition to oral anticoagulant therapy, can influence interpretation of PT test results. Obtaining an accurate patient history and noting specific drug therapies can help in the proper understanding of the potential impact on laboratory test results. The presence of heparin as a contaminant in the patient sample must always be considered when an abnormal result is obtained.

PERFORMANCE CHARACTERISTICS PRECISION

Precision studies were performed to establish Within-Run and Between-Run CV’s for normal control plasma and abnormal control plasma. A single lot number of Thromboplastin LI reagent was used for these studies. Results are shown below

| Photo-optical |  | Within Run | Between Run |
| --- | --- | --- | --- |
| Norm-Trol | n | 20 | 60 |
|  | Mean | 13.7 | 13.7 |
|  | SD | 0.2 | 0.2 |
|  | CV% | 1.7 | 1.3 |
| Ab-Trol 2 | n | 20 | 60 |
|  | Mean | 31.3 | 31.1 |
|  | SD | 0.5 | 0.6 |
|  | CV% | 1.6 | 1.8 |
| Ab-Trol 3 | n | 20 | 60 |
|  | Mean | 66.6 | 68.6 |
|  | SD | 1.5 | 2.1 |
|  | CV% | 2.2 | 3.1 |

Mechanical 		Within Run	 Between Run 20 60

Norm-Trol 	n	20	 60

Mean	14.1 	14.4

SD	0.3 	0.36

CV% 	2.0 	2.5

Ab-Trol 2 	n	20 	60
Mean	32.9 	2.4
SD	0.9 	1.1
CV%	2.8 	3.4

Ab-Trol 3 	n	20 	60
Mean	70.4 	73.5
SD	2.7 	3.4
CV%	3.9 	4.7

COMPARISON

A comparison study was done using the Thromboplastin LI Reagent and another thromboplastin reagent. One hundred five (105) plasma specimens, both normal and abnormal clinical samples were tested with both PT reagents. The linear regression equation and coefficient of determination (r2) of the INR values are reported.

Photo-optical

| N = 105 | r2 = 0.98 | Y = 1.065X – 0.1245 |
| --- | --- | --- |
| Mechanical |  |  |
| N = 105 | r2 = 0.92 | Y = 1.059X – 0.114 |

Y = Thromboplastin LI reagent
X = Reference thromboplastin reagent

References:

1. Quick A.J., The Prothrombin Time in Hemophilia and in Obstructive Jaundice. J. Biol. Chem:109,73-74;1935.

2. Biggs 	R. ed, Human Blood Coagulation Hemostasis and Thrombosis Second Ed. Blackwell Scientific Publications, London 1976.

3. Peterson C.E., Kwaan H.C., Current Concepts of Warfarin Therapy, Arch Intern. Med. 146:581-584, 1986.

4. Loeliger E.A:ICEH/ICTH Recommendations for Reporting Prothrombin Time in Oral Anticoagulant Control, Throm. Haemost. 53:155-156, 1985.

| THROMBOPLASTIN LI Reagent |
| --- |
| Cat. No. |
| Thromboplastin LI Reagent (10 x 5 mL) | 5248 |
| Thromboplastin LI Reagent (10 x 10 mL) | 5249 |
| EQUIPMENT AND SUPPLIES |  |
| Coagulation S.A.R.P. (10 x 1.0 mL) | 5185 |
| Norm-Trol Coagulation Control (10 x 1.0 mL) | 5186 |
| Ab-Trol 1 Coagulation Control (10 x 1.0 mL) | 5187 |
| Ab-Trol 2 Coagulation Control (10 x 1.0 mL) | 5183 |
| Cascade M | 1710 |
| Cascade M-4 | 1711 |

For Sales, Technical and Order Information and Service Assistance, call 800-231-5663 toll free.

Helena Laboratories warrants its products to meet our published specifications and to be free from defects in materials and workmanship. Helena’s liability under this contract or otherwise shall be limited to replacement or refund of any amount not to exceed the purchase price attributable to the goods as to which such claim is made. These alternatives shall be buyer’s exclusive remedies. In no case will Helena Laboratories be liable for consequential damages even if Helena has been advised as to the possibility of such damages. The foregoing warranties are in lieu of all warranties expressed or implied including, but not limited to, the implied warranties of merchantability and fitness

for a particular purpose.

European Community Authorized Representative Helena Laboratories UK, Ltd. Colima Avenue Sunderland Enterprise Park Sunderland SR5 3XB England

Pro. 121

3/07(3)